## Dose Calculation
Small animal orthovoltage dosimetry for XRAD 320
UT SOUTHWESTERN MEDICAL CENTER AT DALLAS
| Calculated by: | Z.Zhang, S.Stojadinovic
Investigator
| Site of irradiation | Whole Body Irradiation |  |  | Date & Time: | 2015-04-06 | 2:00 | PM
1. Instrumentation
| X-ray Unit Model | XRAD 320
| Serial Number | 893
3. Dose at reference condition:
| Calibrated Distance | SSD = | 62 | cm | Tube Potential | kVp = | 250 | kV
| Applicator |  | MC |  | Tube (mA)x(s) | (mA)x(s) = | 15 | mAs
| Field area |  | 20x20 | cm 2 | Added Filtration | F = | 1.65 | mm Al
| Half Value Layer | HVL = | 0.45 | mm Cu | Exposure Time | t = | 1 | min
[5 empty] | End Effect | dt = | 1.1 | s
4. Dose Rate Dw/dt = Dw/(t-dt) [reference applicator] | [5 empty] | DW/(t-dt) = | 1.74 | Gy/min
| Dose Calculation for SSD Setup: |  |  |  | Irradiator Settings
| Applicator Shape |  | Square |  | Added Filtration | F = | 1.65 | mm Al
| Field Size |  | 20x20 | cm2 | Tube Potential | kVp = | 250 | kVp
[5 empty] | Tube (mA)x(s) | (mA)x(s) = | 15 | mAs
| Source to Surface Distance (SSD) |  | 62 | cm
| Output Factor | OF = | 1 |  | Exposure Time | t = | 0.018333 | min
| Prescribed Dose | PD = |  | Gy
| Depth | d = | 15 | mm
| Percent Depth Dose | PDD = | 95.9 | %
| Container Attenuation | TA = | 0.964
|  | t-dt = | 0

## other_Data (hidden)
SSD | Attenuation | Tube (mA)x(s)
23.5 | 1 | 1
23.6 | x | 2
23.7 |  | 3
23.8 |  | 4
23.9 |  | 5
24 |  | 6
24.1 |  | 7
24.2 |  | 8
24.3 |  | 9
24.4 |  | 10
24.5 |  | 11
24.6 |  | 12
24.7 |  | 13
24.8
24.9
25
25.1
25.2
25.3
25.4
25.5
25.6
25.7
25.8
25.9
26
26.1
26.2
26.3
26.4
26.5
26.6
26.7
26.8
26.9
27
27.1
27.2
27.3
27.4
27.5
27.6
27.7
27.8
27.9
28
28.1
28.2
28.3
28.4
28.5
28.6
28.7
28.8
28.9
29
29.1
29.2
29.3
29.4
29.5
29.6
29.7
29.8
29.9
30
30.1
30.2
30.3
30.4
30.5
30.6
30.7
30.8
30.9
31
31.1
31.2
31.3
31.4
31.5
31.6
31.7
31.8
31.9
32
32.1
32.2
32.3
32.4
32.5
32.6
32.7
32.8
32.9
33
33.1
33.2
33.3
33.4
33.5
33.6
33.7
33.8
33.9
34
34.1
34.2
34.3
34.4
34.5
34.6
34.7
34.8
34.9
35
35.1
35.2
35.3
35.4
35.5
35.6
35.7
35.8
35.9
36
36.1
36.2
36.3
36.4
36.5
36.6
36.7
36.8
36.9
37

## Sheet1 (hidden)
30.5 | 0 | 0 | 6.164e-19 | 100 | 30.5 | 0 | 6.27e-19 | 100 | 30.5 | 0 | 6.204e-19 | 100 | 30.5 | 0 | 6.367e-19 | 100
30.6 | 0.1 | 1 | 6.246e-19 | 101.3303 | 30.6 | 1 | 6.349e-19 | 101.26 | 30.6 | 1 | 6.095e-19 | 98.2431 | 30.6 | 1 | 6.323e-19 | 99.3089
30.7 | 0.2 | 2 | 6.235e-19 | 101.1518 | 30.7 | 2 | 6.115e-19 | 97.5279 | 30.7 | 2 | 6.077e-19 | 97.9529 | 30.7 | 2 | 6.293e-19 | 98.8378
30.8 | 0.3 | 3 | 5.833e-19 | 94.6301 | 30.8 | 3 | 6.059e-19 | 96.6348 | 30.8 | 3 | 5.93e-19 | 95.5835 | 30.8 | 3 | 6.081e-19 | 95.5081
30.9 | 0.4 | 4 | 5.721e-19 | 92.8131 | 30.9 | 4 | 5.866e-19 | 93.5566 | 30.9 | 4 | 5.83e-19 | 93.9716 | 30.9 | 4 | 6.057e-19 | 95.1311
31 | 0.5 | 5 | 5.625e-19 | 91.2557 | 31 | 5 | 5.816e-19 | 92.7592 | 31 | 5 | 5.611e-19 | 90.4417 | 31 | 5 | 5.818e-19 | 91.3774
31.1 | 0.6 | 6 | 5.6e-19 | 90.8501 | 31.1 | 6 | 5.575e-19 | 88.9155 | 31.1 | 6 | 5.443e-19 | 87.7337 | 31.1 | 6 | 5.603e-19 | 88.0006
31.2 | 0.7 | 7 | 5.436e-19 | 88.1895 | 31.2 | 7 | 5.402e-19 | 86.1563 | 31.2 | 7 | 5.43e-19 | 87.5242 | 31.2 | 7 | 5.514e-19 | 86.6028
31.3 | 0.8 | 8 | 5.052e-19 | 81.9598 | 31.3 | 8 | 5.138e-19 | 81.9458 | 31.3 | 8 | 5.355e-19 | 86.3153 | 31.3 | 8 | 5.452e-19 | 85.629
31.4 | 0.9 | 9 | 5.517e-19 | 89.5036 | 31.4 | 9 | 5.213e-19 | 83.1419 | 31.4 | 9 | 5.011e-19 | 80.7705 | 31.4 | 9 | 5.197e-19 | 81.624
31.5 | 1 | 10 | 5.082e-19 | 82.4465 | 31.5 | 10 | 4.877e-19 | 77.7831 | 31.5 | 10 | 4.956e-19 | 79.8839 | 31.5 | 10 | 5.049e-19 | 79.2995
31.6 | 1.1 | 11 | 4.762e-19 | 77.255 | 31.6 | 11 | 4.999e-19 | 79.7289 | 31.6 | 11 | 4.925e-19 | 79.3843 | 31.7 | 12 | 4.858e-19 | 76.2997
31.7 | 1.2 | 12 | 4.918e-19 | 79.7859 | 31.7 | 12 | 4.869e-19 | 77.6555 | 31.7 | 12 | 4.55e-19 | 73.3398 | 31.9 | 14 | 4.614e-19 | 72.4674
31.8 | 1.3 | 13 | 4.62e-19 | 74.9513 | 31.8 | 13 | 4.61e-19 | 73.5247 | 31.8 | 13 | 4.604e-19 | 74.2102 | 32.1 | 16 | 4.368e-19 | 68.6037
31.9 | 1.4 | 14 | 4.646e-19 | 75.3731 | 31.9 | 14 | 4.571e-19 | 72.9027 | 31.9 | 14 | 4.544e-19 | 73.2431 | 32.3 | 18 | 4.257e-19 | 66.8604
32 | 1.5 | 15 | 4.378e-19 | 71.0253 | 32 | 15 | 4.43e-19 | 70.6539 | 32 | 15 | 4.314e-19 | 69.5358 | 32.5 | 20 | 3.968e-19 | 62.3213
32.1 | 1.6 | 16 | 4.33e-19 | 70.2466 | 32.1 | 16 | 4.286e-19 | 68.3573 | 32.1 | 16 | 4.268e-19 | 68.7943 | 32.7 | 22 | 3.875e-19 | 60.8607
32.2 | 1.7 | 17 | 4.346e-19 | 70.5062 | 32.2 | 17 | 4.261e-19 | 67.9585 | 32.2 | 17 | 4.176e-19 | 67.3114 | 32.9 | 24 | 3.597e-19 | 56.4944
32.3 | 1.8 | 18 | 4.3e-19 | 69.7599 | 32.3 | 18 | 4.211e-19 | 67.1611 | 32.3 | 18 | 4.037e-19 | 65.0709 | 33.1 | 26 | 3.457e-19 | 54.2956
32.4 | 1.9 | 19 | 4.135e-19 | 67.0831 | 32.4 | 19 | 4.142e-19 | 66.0606 | 32.4 | 19 | 3.995e-19 | 64.3939 | 33.3 | 28 | 3.283e-19 | 51.5627
32.5 | 2 | 20 | 3.744e-19 | 60.7398 | 32.5 | 20 | 3.948e-19 | 62.9665 | 32.5 | 20 | 3.885e-19 | 62.6209 | 33.5 | 30 | 2.99e-19 | 46.9609
32.7 | 2.2 | 22 | 3.663e-19 | 59.4257 | 32.7 | 22 | 3.673e-19 | 58.5805 | 32.7 | 22 | 3.693e-19 | 59.5261 | 34 | 35 | 2.704e-19 | 42.469
32.9 | 2.4 | 24 | 3.616e-19 | 58.6632 | 32.9 | 24 | 3.526e-19 | 56.236 | 32.9 | 24 | 3.464e-19 | 55.8349 | 34.5 | 40 | 2.35e-19 | 36.9091
33.1 | 2.6 | 26 | 3.369e-19 | 54.6561 | 33.1 | 26 | 3.263e-19 | 52.0415 | 33.1 | 26 | 3.263e-19 | 52.5951 | 35 | 45 | 2.116e-19 | 33.2339
33.3 | 2.8 | 28 | 3.248e-19 | 52.6931 | 33.3 | 28 | 3.177e-19 | 50.6699 | 33.3 | 28 | 3.153e-19 | 50.8221 | 35.5 | 50 | 1.872e-19 | 29.4016
33.5 | 3 | 30 | 3.127e-19 | 50.73 | 33.5 | 30 | 3.058e-19 | 48.7719 | 33.5 | 30 | 2.962e-19 | 47.7434 | 36 | 55 | 1.639e-19 | 25.7421
33.7 | 3.2 | 32 | 2.916e-19 | 47.3069 | 33.7 | 32 | 2.91e-19 | 46.4115 | 33.7 | 32 | 2.86e-19 | 46.0993 | 36.5 | 60 | 1.465e-19 | 23.0093
33.9 | 3.4 | 34 | 2.841e-19 | 46.0902 | 33.9 | 34 | 2.733e-19 | 43.5885 | 33.9 | 34 | 2.784e-19 | 44.8743 | 37 | 65 | 1.299e-19 | 20.4021
34.1 | 3.6 | 36 | 2.622e-19 | 42.5373 | 34.1 | 36 | 2.575e-19 | 41.0686 | 34.1 | 36 | 2.584e-19 | 41.6505 | 37.5 | 70 | 1.173e-19 | 18.4231
34.3 | 3.8 | 38 | 2.594e-19 | 42.0831 | 34.3 | 38 | 2.524e-19 | 40.2552 | 34.3 | 38 | 2.485e-19 | 40.0548 | 38 | 75 | 1.035e-19 | 16.2557
34.5 | 4 | 40 | 2.351e-19 | 38.1408 | 34.5 | 40 | 2.375e-19 | 37.8788 | 34.5 | 40 | 2.386e-19 | 38.4591
34.7 | 4.2 | 42 | 2.189e-19 | 35.5127 | 34.7 | 42 | 2.226e-19 | 35.5024 | 34.7 | 42 | 2.175e-19 | 35.058
34.9 | 4.4 | 44 | 2.075e-19 | 33.6632 | 34.9 | 44 | 2.148e-19 | 34.2584 | 34.9 | 44 | 2.113e-19 | 34.0587
35.1 | 4.6 | 46 | 2.126e-19 | 34.4906 | 35.1 | 46 | 2.053e-19 | 32.7432 | 35.1 | 46 | 2.059e-19 | 33.1883
35.3 | 4.8 | 48 | 1.769e-19 | 28.6989 | 35.3 | 48 | 1.959e-19 | 31.244 | 35.3 | 48 | 1.924e-19 | 31.0123
35.5 | 5 | 50 | 1.805e-19 | 29.2829 | 35.5 | 50 | 1.831e-19 | 29.2026 | 35.5 | 50 | 1.841e-19 | 29.6744
35.7 | 5.2 | 52 | 1.747e-19 | 28.342 | 35.7 | 52 | 1.761e-19 | 28.0861 | 35.7 | 52 | 1.761e-19 | 28.3849
35.9 | 5.4 | 54 | 1.774e-19 | 28.78 | 35.9 | 54 | 1.654e-19 | 26.3796 | 35.9 | 54 | 1.62e-19 | 26.1122
36.1 | 5.6 | 56 | 1.698e-19 | 27.547 | 36.1 | 56 | 1.649e-19 | 26.2998 | 36.1 | 56 | 1.557e-19 | 25.0967
36.3 | 5.8 | 58 | 1.528e-19 | 24.7891 | 36.3 | 58 | 1.555e-19 | 24.8006 | 36.3 | 58 | 1.503e-19 | 24.2263
36.5 | 6 | 60 | 1.479e-19 | 23.9942 | 36.5 | 60 | 1.523e-19 | 24.2903 | 36.5 | 60 | 1.419e-19 | 22.8723
36.8 | 6.3 | 63 | 1.308e-19 | 21.22 | 36.8 | 63 | 1.405e-19 | 22.4083 | 36.8 | 63 | 1.307e-19 | 21.0671
37.1 | 6.6 | 66 | 1.264e-19 | 20.5062 | 37.1 | 66 | 1.227e-19 | 19.5694 | 37.1 | 66 | 1.249e-19 | 20.1322
37.4 | 6.9 | 69 | 1.173e-19 | 19.0299 | 37.4 | 69 | 1.151e-19 | 18.3573 | 37.4 | 69 | 1.166e-19 | 18.7943
37.7 | 7.2 | 72 | 1.103e-19 | 17.8942 | 37.7 | 72 | 1.099e-19 | 17.5279 | 37.7 | 72 | 1.091e-19 | 17.5854
38 | 7.5 | 75 | 1.041e-19 | 16.8884 | 38 | 75 | 1.056e-19 | 16.8421 | 38 | 75 | 1.026e-19 | 16.5377
38.3 | 7.8 | 78 | 9.858e-20 | 15.9929 | 38.3 | 78 | 9.973e-20 | 15.9059 | 38.3 | 78 | 9.36e-20 | 15.087
38.6 | 8.1 | 81 | 8.945e-20 | 14.5117 | 38.6 | 81 | 9.047e-20 | 14.429 | 38.6 | 81 | 8.887e-20 | 14.3246
38.9 | 8.4 | 84 | 8.292e-20 | 13.4523 | 38.9 | 84 | 8.331e-20 | 13.2871 | 38.9 | 84 | 8.138e-20 | 13.1173
39.2 | 8.7 | 87 | 8.232e-20 | 13.355 | 39.2 | 87 | 7.867e-20 | 12.547 | 39.2 | 87 | 7.637e-20 | 12.3098
39.5 | 9 | 90 | 7.429e-20 | 12.0522 | 39.5 | 90 | 7.211e-20 | 11.5008 | 39.5 | 90 | 7.262e-20 | 11.7054
39.8 | 9.3 | 93 | 6.877e-20 | 11.1567 | 39.8 | 93 | 6.703e-20 | 10.6906 | 39.8 | 93 | 6.706e-20 | 10.8092
40.1 | 9.6 | 96 | 6.23e-20 | 10.1071 | 40.1 | 96 | 6.403e-20 | 10.2121 | 40.1 | 96 | 6.282e-20 | 10.1257
40.4 | 9.9 | 99 | 5.808e-20 | 9.4225 | 40.4 | 99 | 5.851e-20 | 9.3317 | 40.4 | 99 | 5.723e-20 | 9.2247
40.7 | 10.2 | 102 | 5.916e-20 | 9.5977 | 40.7 | 102 | 5.894e-20 | 9.4003 | 40.7 | 102 | 5.241e-20 | 8.4478
41 | 10.5 | 105 | 5.374e-20 | 8.7184 | 41 | 105 | 5.355e-20 | 8.5407 | 41 | 105 | 5.137e-20 | 8.2801
41.3 | 10.8 | 108 | 5.337e-20 | 8.6583 | 41.3 | 108 | 5.057e-20 | 8.0654 | 41.3 | 108 | 4.786e-20 | 7.7144
41.6 | 11.1 | 111 | 5.156e-20 | 8.3647 | 41.6 | 111 | 4.991e-20 | 7.9601 | 41.6 | 111 | 4.351e-20 | 7.0132
41.9 | 11.4 | 114 | 4.192e-20 | 6.8008 | 41.9 | 114 | 4.234e-20 | 6.7528 | 41.9 | 114 | 3.985e-20 | 6.4233
42.2 | 11.7 | 117 | 3.961e-20 | 6.426 | 42.2 | 117 | 3.966e-20 | 6.3254 | 42.2 | 117 | 3.764e-20 | 6.0671